PT. SINKO PRIMA ALLOY

Alamat		: Jl. Tambak Osowilangun Permai No. 61,

Pergudangan Osowilangun Permai Blok E7-E8, Surabaya-Indonesia (60191)

Telepon		: 031-7482816

Fax. 			: 031-7482815

Aftersale (WA)	: 0821-4281-7085

Email			: aftersales@elitech.co.id

sinkoprima@gmail.com

Website 		: www.elitech.id

SPA-PP/PROD-170.12 Februari 2025.01

Contents

Preface	3

Section 1  Safety Guidance and Symbols	5

1.1 General Safety	5

1.2 Warnings	6

1.3 Cautions	8

1.4 Symbols	8

Section 2  Introduction	10

2.1 Intended Use	10

2.2 Product Features	10

Section 3  Appearance	11

3.1 Appearance	11

3.2 Display	12

3.3 Controlling Keys	12

3.4 Indicating Lights	13

3.5 FHR Work Modes	13

3.6 Display Modes	14

3.7 Parameters Explanations	14

Section 4  Basic Operation	16

4.1 Preparing to Use	16

4.2 Using Battery	16

4.3 Operating Probes	16

4.4 Turning on the Device	16

4.5 Setting Parameter and Working Operation	17

4.6 Assembling Paper	17

4.7 Freezing, Storing and Playing back	17

4.8 Turn off the Device	18

Section 5  Inspecting and Recording…………………………………….….19

5.1 Inspecting FHR	19

5.2 FHR Monitoring	19

5.3 TOCO Monitoring	20

Section 6  Cleaning and Disinfecting	21

6.1 Cleaning	21

6.2 Disinfecting	21

Section 7  Maintenance and Troubleshooting	22

7.1 Maintenance	22

7.2 Troubleshooting	22

Section 8  Warranty and After-sale Service	24

8.1 Warranty	24

8.2 After-sale Service	24

Section 9  Product Specifications	25

# Preface

0.1 Copyright

© All copyright is reserved by manufacturer.

0.2 Attention

This operation manual is written and compiled in accordance with the council directive MDD93/42/EEC for medical devices and harmonized standards.

For any modifications and upgrades, the information in this document are subject to change without notice.

This operation manual includes special documents which are under protection of copyright law.

All rights reserved. Without written announcement from our company, the user manual should not be transferred, copied or translated into other languages.

Our company assumes no responsibility for incidental or consequential damage in connection with the furnishings, performance or use of this material.

0.3 Manufacturer’s Responsibility

The manufacturer only consider itself responsibility for any effects on safety, reliability and performance of the equipment if:

Assembly operations, extensions, readjustments, modifications or reparation are carried out by persons authorized by manufacturer, and the electrical installation of the relevant room compiles with national standard, and the document is used in accordance with the instruction for use.

If necessary, we can provide necessary circuit diagram and other documents to help qualified technicians to maintain and repair the device.

0.4 Warranty
The unit cannot be repaired by users. All services must be done by the engineers approved by manufacturer. We warrant that each product is free from defects in labor and materials and shall conform to its product specifications as defined in the user documentation. If the product doesn't function as warranted during the warranty period, we will repair or replace it without charge. Misuse, improper maintenance may void the warranty.

0.5 Explanation

: The label refers to follow operation manual.

NOTE: Provides useful information of a function or a procedure.

0.6 Warnings

WARNING: The label advises against certain action or situation that could result in personal injury or death.

WARNING: The label advises against actions or situations that could damage equipment, produce inaccurate data, or invalidate a procedure.

Section 1  Safety Guidance and Symbols

### 1.1 General Safety

Before using the device, carefully examine the fetal heart Doppler (hereinafter called device) and the accessories to ensure the main unit and accessories do not have any visible damage evidence that may affect patient safety and device performance.

The device is used for fetal heart rate test, monitoring fetal heart rate and TOCO, not intended for any treatment. If the test result is useless, please try to test by other clinical test methods.

Installment, adjustment, maintains and reparation can be carried out only by qualified or authorized personnel from manufacturer.

It is forbidden to copy or translate any part of the content of the manual to other languages without getting written permission from manufacturer.

Please observe the WARNING, CAUTION and NOTE to avoid possible injury.

WARNING: The device uses very low power ultrasound Doppler. It is confirmed by design calculation, laboratory test, clinical test and clinical application the device Doppler energy is safe for fetus, pregnant women and other personnel. Even so, it is not appropriate to use the device continuously or with long term.

WARNING: To avoid risk of electric shock, this equipment must only be connected to a supply mains with protective earth.

NOTE：Before using the device, please read this manual carefully and assure to be familiar with the controls, displays, features and operating techniques.

### 1.2 Warnings

WARNING: This device is not explosion-proof and cannot be used in the presence of flammable anesthetics equipment.

WARNING: Do not throw battery into fire as this may explode and cause danger.

WARNING: Do not attempt to recharge normal dry-cell battery, which may leak and cause fire or even explosion.

WARNING: Do not touch signal input or output connector and the patient simultaneously to avoid device damage.

WARNING: Accessory and equipment connected to the analog and digital interfaces must be certified according to the respective IEC standards.

WARNING: The battery must be taken out from the battery compartment if the device will not be used for a long time.

WARNING: The operator does not contact patient when changing battery and opening battery cover.

WARNING: Replacing battery shall only be done outside the patient environment (1.5m away from the patient) by persons from manufacturer or persons authorized by manufacturer.

WARNING: Please use the specialized probe from manufacturer.

WARNING: The device is designed for continuous operation and is ordinary. Do not immerse in any liquid (i.e. not drip or splash-proof).

WARNING: Keep the device clean. Avoid vibration.

WARNING: Do not use high temperature sterilizing process and E-beam or gamma radiation sterilization.

WARNING: The device is not subject to any sources of strong electromagnetic interference, such as radio transmitters, mobile telephones, etc. because of electromagnetic interference.

WARNING: The user must check that the equipment does not have visible damage that may affect patient safety or monitoring capability before use. The recommended inspection interval is once per month or less. If damage is evident, replacement is recommended before use.

WARNING: The device cannot be used when some equipment are used, such as high frequency electrical generator, microwave oven and mobile phone.

WARNING: Do not use the device in the presence of flammable anesthetic mixture with oxygen or other flammable agents.

WARNING: Please stop using the device to deal with properly if the device construction is not integrated, such as the battery cover is lost.

WARNING: The device is feasible for relatives such as operator, responsible organization and environmental restriction, which does not need special skills, training and knowledge.

WARNING: This operation manual is written at a level without special education, training and other needs of individual for whom they are intended. For using in hospital and clinic, the operator needs to have relative qualification certificate such as nurse certificate. For using at home, when the test data are used to diagnose, it should be analyzed by specialized medical working personnel.

WARNING: This minimum qualifications of service personnel is to be familiar with our device operation and service technique.

WARNING: For the women to be sensitive to gel, please change gel to water or oil.

WARNING: If using unmatched earphone, it will affect sound effect.

WARNING: After packing, the machine can be sent by air, sea, land. When transporting, the machine cannot be stored in the open.

WARNING: The following safety checks should be performed once every half an year or as specified in the institution’s test and inspection protocol by a qualified person who has adequate training, knowledge, and practical experience to perform these tests.

- Inspect the equipment for mechanical and functional damage.

- Inspect the safety relevant labels for legibility.

- Verify that the device functions properly as described in the instructions for use.

- Test the patient leakage current according to IEC 60601-1: limits less than 100uA (B).

The leakage current should never exceed the limit. The data should be recorded in an equipment log. If the device is not functioning properly or fails any of the above tests, the device has to be repaired.

### 1.3 Cautions

CAUTION: The battery must be properly disposed according to local regulation after use.

CAUTION: The device shall only be used when the battery cover is closed.

CAUTION: Battery must be stored in cool and dry place.

CAUTION: If use rechargeable battery, to insure capability and life, please fully charge batteries before first use, normally, batteries must be continuously charged over 4 hours or charged according to the guidance displayed on the battery.

CAUTION: Please do not set anode and cathode of the battery wrongly.

CAUTION: The information in this manual are subject to modify without notice.

The device usage life is 6 years, after which please treat according to relative regulations.

CAUTION: Do not use strong solvent (for example, acetone) and abrasion material to clean the device.

CAUTION: Do not allow any liquid to enter the device, and do not immerse any parts of the device into any liquids.

CAUTION: Please choose the accessories and expendable authorized by manufacturer.

CAUTION: Never try to sterilize the probe or equipment by low temperature steam or other methods.

### 1.4 Symbols

All the symbols are described as follows.

| Symbol | Explanation |
| --- | --- |
|  | Type BF |
|  | Attention, refer to the accompanying documents |
|  | Headphone socket |
|  | Power ON/OFF |
|  | Fetal Heart Rate symbol |
|  | This item is compliant with Medical Device Directive 93/42/EEC |
| P/N | Part number |
| S/N | Serial number |
|  | Date of manufacture |
|  | Manufacturer |
|  | Authorized representative in the European community |

# Section 2  Introduction

### 2.1 Intended Use

- 2MHz FHR probe: detect fetal heart rate from 10-12weeks; Position for the fetus before assembling ultrasound monitoring probe.

- 1MHz ultrasound monitoring probe: continuously monitor FHR.

- Stimulator: stimulate to wake up the fetus if the fetus is sleeping by vibration when continuous monitoring FHR by 1Mhz ultrasound probe.

- TOCO probe: continuously monitor TOCO (Urine Contraction Pressure).

### 2.2 Product Features

Different models have different features as follows.

| Function | FD-800T | FD-800P | FD-800B |
| --- | --- | --- | --- |
| Display | Color | Color | Color |
| Touch screen | Yes | Yes | Yes |
| Code switch | Yes | Yes | Yes |
| Audio/light alarm | Yes | Yes | Yes |
| Printing system | Yes | Yes | Yes |
| FHR detect | Yes | Yes | Yes |
| FHR monitor | Yes | Yes | Yes |
| TOCO monitor | Yes | No | No |
| Stimulator | Yes | Yes | Yes |
| SpO2 | No | Yes | No |
| PC software | Option | Option | Option |
| Carry case | Option | Option | Option |
| Trolley | Option | Option | Option |

# Section 3  Appearance

### 3.1 Appearance

Front

Back               Bottom

Remark:

1) Main body        		2) Printer cover         	3) AC light

4) Power On/Off     		5) Power light          	6) Printing key

7) Alarm On/Off Key  		8) Freezing key  　     	9) Stimulator key

10) Zero key        		11) Probe Change key    	12) Code switch

13) 1MHz probe     		14) TOCO probe         	15) 2Mhz probe

16) 2Mhz probe port      17) 1Mhz probe			18) Mark port

19) TOCO probe port		20) USB port     		21) Power switch

22) DC port             23) Screw hole

### 3.2 Display

3.2.1 Working Display

Digit+Curve mode      Curve mode       Digit mode

3.2.2Menu Displays

SONOTRAX PRO TFT (TOCO) menu

### 3.3 Controlling Keys

In 3.1 main unit configuration, these keys as follows.

3.3.1Power ON/OFF

Press  key once to turn on the device.

When the device is working, press  key for 3 seconds to turn off  the device.

3.3.2 Printing Key

Print  key to start printing till the setting printing time.

When printing, press  key to stop printing.

3.3.3 Alarm ON/OFF Key

When existing alarm, press  key to stop alarm sound temporarily. After the delay time, it will sound again if the alarm condition still exists.

3.3.4 Freezing Key

Press  key to freeze contents on LCD will be freeze. This will keep till another measurement starting, changing the mode or pressing this key again.

3.3.5 Stimulator Key

If the fetus is sleeping when continuous monitoring FHR by 1Mhz ultrasound probe, Press  key to wake up the fetus by vibration.

3.3.6 Zero Key

Press  key to set the TOCO to 0 value.

3.3.7 Probe Change key

Press  key to exchange the probe between 2MHz probe and 1 MHz ultrasound probe.

3.3.8 Code Switch

Code switch is used to select and set parameter and function. Rotate the switch to suitable place to select and press to confirm.

### 3.4 Indicating Lights

There are two lights with code (3) and (5) in main unit configuration.

3.4.1 Power Indicator Light

When turning on the device, the light is bright all the time.

3.4.2 Charge Indicator Light

When charging, it is orange color. After finishing charging, it is green.

### 3.5 FHR Work Modes

For any FHR work modes, the FHR is automatically displayed on LCD. There are these work modes as follows.

3.5.1 Real-time Mode

In this mode, the heart symbol on LCD will flash, and real-time FHR is displayed on LCD simultaneously. You can record or stop recording by pressing  key once.

3.5.2 Averaged Mode

This mode is used to get more stable heart rate values. The LCD displays the flashing heart symbol when displaying FHR. The FHR value is averaged and displayed during the setting time.

3.5.3 Manual Mode

This mode is used for FHR is not enough to display for Doppler but can be audible.

Press Power key once, start to count. The LCD displays the flashing heart shape symbol and “---”. Press Power key again on the tenth count (after 9 beat internals). The device automatically calculates the derived FHR averaged over the 10 beat periods and display the FHR value. This rate value will keep till another measurement starts or the mode is changed.

3.5.4 Demo Mode

This mode is to display the data which is stored into device in advance.

### 3.6 Display Modes

There are these display modes: curve mode, digit mode, curve + digit mode.

### 3.7 Parameters Explanations

| Sub-menu | Parameter | Explanation | Selectable value | Default value |
| --- | --- | --- | --- | --- |
| FHR setup | FHR Volume | FHR volume | 0-7 | 0 |
|  | Color | FHR color | CYAN, GREEN, PURPLE | GREEN |
|  | Print speed | Print speed | 1, 2, 3mm/min | 3mm/min |
|  | Print time | Automatic print time | 00-60s | 00 |
| TOCO setup | Wave thick | The size of TOCO wave | 1,2 | 1 |
|  | TOCO color | TOCO color | Green, Cyan, purple, yellow, white, orange | GREEN |
| SpO2 setup | Scan speed | Scan speed per display page | 12.5mm/s,25mm/s | 25mm/s |
|  | SpO2 color | Color of SpO2 digit | Green, Cyan, purple, yellow, white, orange | Green |
|  | PR color | Color of PR digit | Green, Cyan, purple, yellow, white, orange | Green |
|  | Filling | Filling curve ON or OFF | ON or OFF | ON |
|  | Alarm | Filling curve On or OFF | ON or OFF | ON |
| Alarm setup | Alarm | Alarm ON or OFF | ON, OFF | ON |
|  | Upper limit | FHR Alarm upper limit | 50-240bpm | 160bpm |
|  | Lower limit | FHR Alarm lower limit | 50-240bpm | 120bpm |
|  | Delay time | Alarm delay time | 5-20S | 10S |
| Factory set | Factory set | Factory default set | YES,NO | NO |
| Display mode | Only curve | Display curve and other information |  | YES |
|  | Only number | Display curve FHR number and other information |  | / |
|  | Number &info | Display FHR, parameters, etc. |  | / |
| Time setup | Year | Year | 00-99 | 10 |
|  | Month | Month | 1-12 | 01 |
|  | Day | Day | 1-31 | 01 |
|  | Hour | Hour | 00-23 | 12 |
|  | Minute | Minute | 00-59 | 04 |
| Work mode | Real time mode | Real time mode |  | YES |
|  | Average mode | Average mode |  | / |
|  | Manual mode | Manual mode |  | / |
|  | Demo mode | Demo mode | ON,OFF | OFF |
| System set | Language | Language | ENGLISH,CHINESE | ENGLISH |
|  | Touch panel adjustment | Touch panel adjustment | / | / |

# Section 4  Basic Operation

### 4.1 Preparing to Use

Be carefully check if the device has any damnification and if the accessories are integrated. If so, please stop using the device and contact manufacturer or local distributor.

Keep the package for possible future transportation and storage.

Be carefully take out the probe from the probe slot at the side of the main unit. Please pay attention not to pull the probe forcibly when operating.

### 4.2 Using Battery

4.2.1 Taking out Battery

Take away the battery cover and take out the battery.

4.2.2 Placing the Battery

Insert the plug of battery into the battery socket and put the battery into the compartment with the wire inward.

4.2.3 Closing the Battery Compartment

Put the battery cover, screw the screw.

### 4.3 Operating Probes

4.3.1 Taking out and Inspecting the Probe

Before using the device, take out and inspect the probe carefully.

4.3.2 Replacing Probe

The probe is changeable. Before doing it, please turn off the device.

Press the spring plate of crystal head, pull out the plug of the probe from its socket, then connect the probe you need to the socket.

### 4.4 Turning on the Device

Turn on by pressing the power  key once, the working indicator light is bright and the equipment can start to work.

### 4.5 Setting Parameter and Working Operation

- There are three ways to operate the device: code switch, key operation and touch screen.

- Press code switch to enter into menu to set working parameters. Press up key and down key to select sub menu and parameter. Press  confirmation key to confirm the selected sub menu and parameter, and exit setting state to begin to work.

The setting parameters are automatically saved.

When working, press  key to freeze the display on screen.

When working, press  up key and down key to increase or reduce the FHR volume.

When working, press alarm key to turn or turn off the alarm voice.

Press  print key to start or stop printing.

At the bottom of screen, the working state is displayed.

### 4.6 Assembling Paper

Press the sliding key at the side of device, open the paper cassette, take out the paper shaft, assemble the paper on the shaft, place the shaft with paper into the paper cassette, and close the paper cassette cover.

The recording paper should be thermal recording paper with 50mm width and without grid.

### 4.7 Freezing, Storing and Playing back

When working, the working data are automatically stored for every 2 minutes within 24 hours. After 24 hours, the data are automatically renewed by time sequence.

By pressing  key to freeze the working data and display the following state.

The playback time can be set by pressing , they are 2minutes/page,1 hour/page, 8hours/page and 24hours/page.

Press key the change to last or next page.

Press  to turn on/off freeze.

### 4.8 Turn off the Device

When the device is power on and being work state, press the   key again for 3 seconds, the device is power off and the working indicator light is off.

NOTE: The device will automatically turn off with 3 minutes if it is not used.
Section 5  Inspecting and Recording

### 5.1 Inspecting FHR

- Insert the FHR probe plug into the socket.

- Daub ultrasonic gel on the probe faceplate to decrease noise and boost up test result.

- Finding the Position of Fetus

Firstly, please feel the fetal position by hand. Place the working faceplate on abdomen with feasible tight contact. Adjust the probe position to obtain an optimum audio signal. Usually, the heart position of the fetus with small gestational age is on the 1/3 of the line from the umbilicus to symphysis pubis, and with increasing of the gestational age the position will shift upward and deviate to the left or right slightly.

- FHR Inspecting

Finish setting mode and parameters, press power On/Off key to start work, press power On/Off key again to stop working.

NOTE: Do not measure FHR unless audible and identifiable fetal sound has been heard, usually it needs about 5 seconds.

NOTE: The normal value of the fetal heart rate is 120-160bpm. 100-120bpm and 160-180bpm are the critical values which should be paid some attention to. And lower than 100bpm and more than 180bpm are danger values which should be paid more attention to.

### 5.2 FHR Monitoring

- Insert the ultrasound transducer plug into the socket.

- Find the fetal position by palpation or by FHR detecting.

- Daub gel at the ultrasound transducer working face.

- Place the transducer on the abdomen and keep the gel to contact tightly between the skin and the transducer. The best signal will be achieved by positioning the transducer over the fetus’s upper back over the left scapula. Adjust the position slowly until the best signal is obtained.

- When the fetal heartbeat is heard, keep the transducer in position and move one end of the belt into position.

- The fetal heart rate value will be shown on the screen. At the same time, the ultrasound wave will be drawn on the screen.

- During the monitoring, if the FHR can be heard without steadily sound of the fetal heart, it may not proper positions. So move it slowly until the proper position is found. But if it is not found, the doctor should do other examination, to observe if the fetus is normal.

### 5.3 TOCO Monitoring

- Insert the TOCO transducer plug into the socket.

- Position the TOCO transducer over the fundus of the uterus.

- Put the transducer in suitable position, move on end of the belt into position over the transducer so that the belt keep tension and transducer is at suitable position.

- Zero can be set more quickly by pressing the  button on the front panel. The TOCO display should respond to contractions and maternal breath.

NOTE: The display of external pressure is shown as a percentage % of full scale. The uterine activity reading at this point should be greater than 30 units and less than 90 units. If the reading falls outside this range, the belt may be too tight or too loose.

NOTE: Do not use coupling gel on the TOCO transducer contact area.

# Section 6  Cleaning and Disinfecting

### 6.1 Cleaning

Before cleaning the device, switch off the device and take out the battery from main unit.

Keep the outside surface of the device clean and free of dust and dirt, clean exterior surface with a dry, soft cloth. If necessary, clean the chassis with a soft cloth soaked in a solution of soap, or water and wipe dry with a clean cloth immediately.

Wipe the probe with soft cloth to remove any remaining ultrasound gel. Clean with soap and water only.

NOTE: Wipe the surface of probe with 70% ethanol, self-air dry, or clean with a clean, dry cloth.

### 6.2 Disinfecting

Clean the equipment case, probe, etc. as above, and then wipe the probe with an alcohol impregnated wipe (70% ethanol).

Wipe the probe with a clean, dry cloth to remove any remaining moisture.

NOTE: Never try to sterilize the probe or equipment by low temperature steam or other methods.

# Section 7  Maintenance and Troubleshooting

### 7.1 Maintenance

The device is precision equipment, and the probe acoustic surface is frangible, you need to handle the device especially probe with enough care.

Gel and dirty dunghill must be wiped from the probe after using. These precautions will prolong the life of the unit and keep the examination precision.

Before using, the user must check that the equipment does not have visible evidence of damage that may affect patient safety or device capability. The recommended inspection interval is once per week. If damage is evident, reparation is recommended before use.

The equipment should undergo periodic safety testing to insure proper patient isolation from leakage currents. This should include leakage current measurement. The recommended testing interval is once every two years or as specified in the institution’s test and inspection protocol.

The accuracy of FHR is controlled by the device and cannot be adjusted by user. If the FHR result is distrustful, please use other method such as stethoscope to verify immediately or contact local distributor or manufacturer to get help.

There are some optional probes. The probes can be replaced by users.

The replacing of fuse should buy by persons from PT Sinko Prima Alloy or persons authorized by PT Sinko Prima Alloy.

### 7.2 Troubleshooting

When using, if it appears following problems, please treat by following instruction. If fail to treat, please contact local distributor or manufacturer.

7.2.1 No Sound

Main reasons: (1) The battery volume is serious insufficient; (2) The equipment is damaged; (3) The battery connection line is damaged.

Solution: (1) Charge the battery or change battery; (2) Inspect the device; (3) Contact dealer or manufacturer.

7.2.2 Weak Sound

Main reasons: (1) Voice volume is too low; (2) Battery volume is too low; (3) Without sufficient gel.

Solutions:(1) Adjust higher voice volume; (2) Change or charge the batteries; (3) Add sufficient gel on probe inspecting surface.

7.2.3 High Noise

Main reasons: (1) Probe is too near from the main unit; (2) Disturbance from the outside signal; (3) Power is low.

Solutions: (1) Keep the probe feasible from main unit; (2) Be away from the outside signals; (3) Change or charge the battery.

7.2.4 Low Sensitivity

Main reasons: (1) Position is incorrect; (2) Without or insufficient gel.

Solutions: (1) Keep the probe at right position; (2) Daub sufficient gel.

# Section 8  Warranty and After-sale Service

### 8.1 Warranty

PT Sinko Prima Alloy obligation under this warranty is limited to repair any part or whole unit upon PT Sinko Prima Alloy’s examination to prove they are within warranty period and range. If the product does not function during the warranty period, we will repair or replace it free of charge.

Limit of warranty:

1) Trouble resulting from misuse, negligence, accidents or transportation.

2) Opening, modification or repair by unauthorized persons from manufacturer.

3) Replacement or remove serial number label or label.

### 8.2 After-sale Service

If you have any questions of use, maintenance, technical specifications or malfunction of device, please contact local distributor or PT Sinko Prima Alloy service department.

# Section 9  Product Specifications

9.1 Product Name: Desktop Fetal Doppler

9.2 Models: SONOTRAX PRO TFT (TOCO)

9.3 Safety Standard: IEC 60601-1:2005

9.4 Classification

- Anti-electroshock Type: Externally powered equipment.

- Anti-electroshock Degree: Type BF equipment

- Harmful Liquid Proof Degree:IPX1，Ordinary equipment (sealed equipment without liquid proof)

- Degree of Safety in Presence of Flammable Gases: Equipment not suitable for use in presence of flammable gases

9.5 Working System: Continuous running equipment

9.6 EMC: Group I Class B

9.7 Physical Characteristic

- Size: 225mm×220mm×162mm

- Weight: About 2.5Kg (including battery)

9.8 Environment

- Working: Temperature: +5～+40℃

Humidity: 25-80％

Atmospheric Pressure: 70～l06kPa

- Transport and Storage:  Temperature: -10～+55℃

Humidity: ≤93％

Atmospheric Pressure: 50～l06kPa

9.9 Display: 65mm×50mm LCD

9.10 Battery Type Recommended: Ni-MH Battery/16.8V

9.11 2MHz FHR Probe

- Measuring Range: 50~240bpm

- Resolution: 1bpm

- Accuracy: ±1bpm

- Output Power: <0.8W

- Nominal Frequency: 2.0MHz

- Working Frequency: 2.0MHz±10%

- P-: <1MPa

- Iob: <20mW/cm2

- Ispta: <100mW/cm2

- Ultrasonic Output Intensity：Isata<10mW/cm2

9.12 Ultrasound probe

- Nominal Frequency: 1.0MHz

- Work Frequency: 1.0MHz±10％

- Negative peak sound pressure : P_＜1MPa

- Output beam intensity : Iob＜20mW/cm

- The average time space peak intensity: Ispta＜10mW/cm

- FHR Rang: 50BPM～210BPM，Allow error: within ±2bpm

- Resolution: 1BPM

- Accuracy : ±2％

9.13 TOCO

- TOCO range: 0～100％

- Resolution: 1％

- Nonlinear error: ＜±8％

- Zero setting way: Manually

9.14 Recommended Coupling Medium

- Stimulation to Skin: No

- Total Germ Quantity: <1000units/g

- Dung Escherichia Coli, Pseudomonas Aeruginosa and Staphy-lococcus Aureus: No

- Acoustic Velocity: 1520-1620m/s

- Acoustic Impedance: 1.5 -1.7x106 Pa.s/m

- Acoustic Attenuation: <0.05dB/(cm.MHz)

- Viscosity: >15Pa.S

- PH Value: 5.5-8

9.15 Material Group: I

9.16 Pollution Degree: II

9.17 Operating Altitude: <2000m

9.18 Overvoltage Class: I

DESKTOP FETAL DOPPLER

SONOTRAX PRO TFT (TOCO)

MANUAL BOOK