# Manual Terstruktur: manual_fox_baby_en.pdf

# 1.2 Panduan Keamanan

such faults. ● Some functional testers or patient simulators can measure 1.4.1 Attention optical obstacles like rubberized fabric, to avoid inaccurate

Owing to the forthcoming renovation, the specific products you the accuracy of the device copied calibration curve, but they Point out conditions or practices that may cause damage to the results.

received may not be totally in accordance with the description of cannot be used to evaluate the device accuracy. device or other properties. ● Excessive ambient light may affect the measured results, such

this User Manual. We would sincerely regret for that. ● When using the device, please keep it away from the ● Before using the device, make sure that it locates in normal as surgical light (especially xenon light sources), bilirubin

Remind that it may cause serious consequences to tester, user or magnetic field. Using the device in an inappropriate ● In order to get a more accurate measurement, it should be etc. In order to prevent interface from ambient light, make

environment. environment may cause interference to the surrounding radio used in a quiet and comfortable environment. sure to place the sensor properly and cover the sensor with

● Explosive hazard - DO NOT use the device in environment with equipment or affect its working. ● When the device is carried from cold environment to warm or opaque material.

inflammable gas such as anesthetic. ● When storing the device, keep it away from children, pets, humid environment, please do not use it immediately, wait ● Frequent movement (active or passive) of the subject or

● DO NOT use the device while examining by MRI or CT, as the and insects to avoid affecting its performance. four hours at least is recommended. severe activity can affect the measured accuracy.

induced current may cause burn. ● Do not place the device in places exposed to direct sunlight, ● If the device is splashed or coagulated by water, please stop ● The SpO2 probe should not be placed on a limb with the blood

sole basis for clinical diagnosis. The device is only used as an splash water, to avoid affecting its performance. ● DO NOT operate the device with sharp things. ● The measured value may be inaccurate during defibrillation

auxiliary means in diagnosis. And it must be used in ● The measured accuracy will be affected by the interference of ● High temperature, high pressure, gas sterilizing, or immersion and in a short period after defibrillation, as it has not

# 4.4 Pembersihan

battery by the personnel not fully trained. Users are not ● This device is not intended for treatment. satisfactory result, please stop using it. ● Please select medical power adapter to charge it, when

permitted to maintain or refit the device by themselves or ● The intended operator of the device may be a patient. ● Data averaging and signal processing have a delay in the connecting the special adapter with the socket, make sure

replacement of the battery. ● Avoid maintaining the device during using. upgrade of SpO2 data values. When the data update period is there is no shelter near the socket and it is easy to plug and

device ceaselessly, especially for the microcirculation The oxygen saturation is the percentage of HbO2 in the total Hb in values will increase, which is arisen from signal degradation, necessary, causes damage.

disturbance users. It is not recommended that the sensor is the blood, so-called the O2 concentration in the blood, it is an low perfusion or other interference, it depends on the PR 1.4.2 Clinical Restriction

used on the same finger for more than 2 hours. important physiological parameter for the respiratory and value. A. As the measure is taken on the basis of arteriole pulse,

● For some special users who need a more careful inspection on circulatory system. A number of diseases relating to respiratory ● The device has 3 years’ service life, date of manufacture: see substantial pulsating blood flow of subject is required. For

the test site, please don’t place the device on the edema or system may cause the decrease of SpO2 in the blood, furthermore, the label a subject with weak pulse due to shock, low ambient/

tender tissue. some other causes such as the malfunction of human body’s self- ● The maximum temperature at the SpO2 probe – tissue body temperature, major bleeding, or use of vascular

● Please do not stare at the red and infrared light emitter (the adjustment, damages during surgery, and the injuries caused by interface should be less than 41OC which is measured by the contracting drug, the SpO2 waveform (PLETH) will

infrared light is invisible) after turning on the device, including some medical checkup would also lead to the difficulty of oxygen temperature tester. decrease. In this case, the measurement will be more

● The device contains silicone, PVC, TPU, TPE, and ABS appear as a consequence, such as vertigo, impotence, vomit etc. screen, please pull out your finger and reinsert it to measure B. The measurement will be influenced by intravascular

materials, whose biocompatibility has been tested in Serious symptoms might bring danger to human’s life. Therefore, again. staining agents (such as indocyanine green or methylene

passed the recommended biocompatibility test. The person to discover the potential danger, and is of great importance in the ● The plethysmography waveform is not normalized, as a signal C. The measured value may be normal seemingly for the

who is allergic to silicone, PVC, TPU, TPE, or ABS cannot use clinical medical field. inadequacy indicator, when it is not smooth and stable, the tester who has anemia or dysfunctional hemoglobin (such

the device. accuracy of the measured value may degrade. When it tends as carboxyhaemoglobin (COHb), methaemoglobin

to be smooth and stable, the measured value read is the (MetHb), and sulfhaemoglobin (SuHb)), but the tester

symptoms. A. The packed device can be transported by ordinary documents

D. Pulse oxygen only has a reference meaning for anemia conveyance or according to transport contract. During Type BF applied Pulse oxygen

and toxic hypoxia, as some severe anemia patients still transportation, avoid strong shock, vibration and splashing part %SpO2 saturation (%)

show better pulse oxygen measured value. with rain or snow, and it cannot be transported mixed with 1. The finger clip Low-power

E. Contraindication: No. toxic, harmful, corrosive materials. falls off (no indication

2. Principle B. The packed device should be stored in room with no corrosive finger

Figure 3 Sketch map for finger placement gases and good ventilation. Temperature: -40°C~+60°C; inserted)

B. Wait a few seconds, the device directly shows Relative Humidity: ≤95%. - - - 2. Probe error

Note: when inserting the finger, the light emitting from the Trouble Possible Reason Solution inadequacy

sensor must be directly irradiated to the side of the indicator

fingernail. 1. T

ro

in

rl

e r is not

1. Please insert the

Serial Number Use-by date

Figure 1 Operating principle Note: during measuring, do not shake the finger and keep inserted. finger properly Recycling USB

quiet, not move. 2. The finger is and measure garbage WEEE

Principle of the Oximeter is as follows: An experience formula of 6.2 Pulse Sound Setting shaking or the again. (2002/96/EC)

data process is established taking use of Lambert Beer Law After turning the device on, the pulse sound is open. Long press the The values patient is 2. Let the patient The sound The pulse

according to Spectrum Absorption Characteristics of Reductive button can close the pulse sound and the pulse sound indication cannot be moving. keep calm. prompt sound

i O n x fr y a h r e e m d o z g o l n o e b s in . O In p s e p ra e t c i t o io n n p T ri e n c c h ip n l o e l o o g f y t h is e a d d e o v p ic t e e d is : i n P h a o c t c o o e r l d e a c n tr c i e c o 6. n 3 a n S d o p u u n l d se p s r o o u m n p d t i n se d t ic ti a n t g io n appears. n st o a r b m ly a . lly or n en o v t i u ro se n d m i e n n t d en ev v i i c ro e n in m n e o n r t m . al IP22 I P n r t o e t r e n c a t t io io n n al Fi O n u g t e r T n h o e t ins fi e n r g te e d r is

with Capacity Pulse Scanning & Recording Technology, so that two a. Sound prompt includes the sound prompt of measure required by the 4. Please contact the Temperature Fragile, handle

b h e u a m m a s n n o a f il d ti i p ff e th re ro n u t g w h a p v e e r l s e p n e g c t t h iv e o f c la li m gh p t s f in c g a e n r -t b y e p e f o s c e u n s s e o d r. T o h n e to n d 90 at % a , ’ s o r g t o h in e g p u b l e s y e o r n a d te t i h s e n o li t m b i e t t s w ( e w e h n e 5 n 0 t b h p e m S p a O nd 2 1 is 2 0 b e b l p o m w , 4. m Th a e n d u e a v l. i ce after sales. l A im tm it o a s t p io h n e ric w Th it is h w ca a r y e u p

m i t n h f e r o o a r u s m u g r h a e t t d io r e n s a ig t a m n c a q e l u n c i t r a e in n d e b l t e e h c r t o o r b u o t g n a h i i c n c e w i d r h c i u b ch i y t s a w a n i p l d l h o m b t e i o c s r s e o h n p o s r w i o t n i c v e e s o s n e o l r e s . m cr e e n e t n , t p h o e w e so r. u nd prompt occurs), the sound prompts of low- w ab o n r o ks rm ally. p lim re i s t s a u ti r o e n

3. Functions b. In the open state of sound prompt, when the measure data 1. T

ra

i n

tt

ry

or

1. P

te

se

ri e

a rge H

lim

um

ita

id

ti

ty

fr

rain

away

almost drained 2. Please install the Recyclable Manufacture

The device away. battery again. date

E F. .

L t

A D

h o

u is

e w

t p

b -

m a

a t

t t

m t

e t

e r

c o

y r

s d

t e

l d t

a d

a i

g a

e f

i i o

o :

c a t

o lo

i n o

l w

n e

- w b

a to tt e w r o y r i k n dication appears when

b v

in

u l

t u t

e o

n s

, t

i a l

l n

g d

li

t t

h e

e r.

S o o

u u

ld

n n d d

p p

r r o o

m m

us

p p

t t

f i

c u

o n

n c

t d i o i

n s

a w p

p i

l e l

a b

r e

s , r

e b

n u

e t

h e

d e

tu

on.

b e 2

i i

n n

s c

t o a r

l r l

e e

c d

ly

’ s

i s 3. P

l c o e

le

c n a

t l

e s

r e .

ic

a ct the

9.

e:

ci

ic

ti

ic

a r

ib

i t

c ontain all the followi ng symbols.

G. Charging function If low-power sound prompt occurs, please charge the battery. malfunction SpO2 [see note 1]

I. PR sound indication: sound prompt for over-limit, low battery Power adapter can be selected to charge for the device. It indicates enters into the Measured range 0% ~ 99%

4. Installation that the device is charging when the indicator is light, the charging energy saving 1. Normal. Accuracy [see note 2] 70% ~ 99%: ±2%;

disappears 2. Low battery. battery.

suddenly. 3. The device 3. Please contact the

works after-sales.

Pulse bar graph Use 75% alcohol to wipe the device enclosure, and use liquid soap 1. The battery is Accuracy ±2bpm or ±2%, whichever is

or clean it with clean and soft cloth. Do not spray any liquid on cannot be fully. battery. Resolution 1 bpm

Figure 2 Appearance and Measurement Interface the device directly, and avoid liquid penetrating into the device. used for full 2. The device 2. Please contact the Accuracy under low Low perfusion 0.4%:

4.2 Connection of USB cable make sure that there is no visible damage that may affect The battery is greater

Insert the micro end of the USB cable into the device, the other patient’s safety and monitoring performance. It is cannot be Light interference Under normal and ambient light

end into computer or power adapter. recommended that the device should be inspected weekly fully conditions, the SpO2 deviation

(optional), one user manual, one lanyard. C. Please charge the battery in time when low battery charging stronger pulse.

list to avoid that the device cannot work normally. D. Recharge the battery soon after over-discharge. The SpO2 0% ~ 99%

C. Software description device should be recharged every three months when it is 8. Symbols PR 30 bpm ~ 249 bpm

Release version: V2.0 not used for some time. It can extend the battery life Symbols Meaning Symbols Meaning Optical sensor [see note 4]

5. Operating following this guidance. Red light Wavelength: about 660 nm

|  | Caution, consult accompanying documents | PRbpm | Pulse rate (bpm) |
| --- | --- | --- | --- |
|  | Type BF applied part | %SpO2 | Pulse oxygen saturation (%) |
| - - - | 1. The finger clip falls off (no finger inserted) 2. Probe error 3. Signal inadequacy indicator |  | Low-power indication |
|  | Serial Number |  | Use-by date |
|  | Recycling garbage WEEE (2002/96/EC) |  | USB |
|  | The sound prompt indication |  | The pulse sound indication |
| IP22 | International Protection | Finger Out | The finger is not inserted |
|  | Temperature limitation |  | Fragile, handle with care |
|  | Atmospheric pressure limitation |  | This way up |
|  | Humidity limitation |  | Keep away from rain |
|  | Recyclable |  | Manufacture date |
|  | Alarm inhibit |  |  |

| 7. Troubleshooting |  |  |
| --- | --- | --- |
| Trouble | Possible Reason | Solution |
| The values cannot be displayed normally or stably. | 1. The finger is not properly inserted. 2. The finger is shaking or the patient is moving. 3. The device is not used in environment required by the manual. 4. The device works abnormally. | 1. Please insert the finger properly and measure again. 2. Let the patient keep calm. 3. Please use the device in normal environment. 4. Please contact the after sales. |
| The device cannot be turned on. | 1. The battery is drained away or almost drained away. 2. The battery is installed incorrectly. 3. The device’s malfunction | 1. Please charge batteries. 2. Please install the battery again. 3. Please contact the local service center. |
| The display disappears suddenly. | 1. The device enters into the energy saving mode. 2. Low battery. 3. The device works abnormally. | 1. Normal. 2. Please charge the battery. 3. Please contact the after-sales. |
| The device cannot be used for full time after charge. | 1. The battery is not charged fully. 2. The device works abnormally. | 1. Please charge the battery. 2. Please contact the after-sales. |
| The battery cannot be fully charged even after 10-hour charging time. | The device works abnormally. | Please contact the after sales. |

| 9. Specification |  |  |  |
| --- | --- | --- | --- |
|  | SpO2 [see note 1] |  |  |
| Display range |  | 0% ~ 99% |  |
| Measured range |  | 0% ~ 99% |  |
| Accuracy [see note 2] |  | 70% ~ 99%: ±2%; 0% ~ 70%: unspecified. |  |
| Resolution |  | 1% |  |
|  | PR |  |  |
| Display range |  | 30 bpm ~ 249 bpm |  |
| Measured range |  | 30 bpm ~ 249 bpm |  |
| Accuracy |  | ±2bpm or ±2%, whichever is greater |  |
| Resolution |  | 1 bpm |  |
| Accuracy under low perfusion [see note 3] |  | Low perfusion 0.4%: SpO2: ±4%, PR: ±2 bpm or ±2%, whichever is greater |  |
| Light interference |  | Under normal and ambient light conditions, the SpO2 deviation ≤1% |  |
| Pulse intensity |  | Continuous bar graph display, the higher display indicates the stronger pulse. |  |
|  | Upper and lower limit of measured values |  |  |
| SpO2 |  | 0% ~ 99% |  |
| PR |  | 30 bpm ~ 249 bpm |  |
|  | Optical sensor [see note 4] |  |  |
| Red light |  | Wavelength: about 660 nm |  |

| 8. Symbols |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Symbols |  |  | Meaning |  |  | Symbols |  |  | Meaning |  |

Optical output power: <6.65 mW

Infrared light Wavelength: about 905 nm

Optical output power: <6.75 mW

Safety class Internally powered equipment,

type BP applied part

International protection IP22

Working voltage DC 3.6V ~ 4.2V

Working current ≤ 80 mA

Power supply A rechargeable lithium battery

(3.7 V) (The red wire on the

battery denotes anode, the

black wire on the battery

denotes cathode).

Battery life Charge and discharge: no less

than 500 times.

Output current: 1000 mA

Weight About 35 g (including a lithium

battery)

Appendix 1

State Sound prompt Sound prompt

condition delay signal generation

delay

Low voltage 1s 20ms

sound prompt

SpO2 sound 330ms 20ms

prompt

Pulse rate sound 330ms 20ms

prompt

Probe error 16ms 20ms

sound prompt

|  |  | Optical output power: <6.65 mW |  |
| --- | --- | --- | --- |
| Infrared light |  | Wavelength: about 905 nm Optical output power: <6.75 mW |  |
| Safety class |  | Internally powered equipment, type BP applied part |  |
| International protection |  | IP22 |  |
| Working voltage |  | DC 3.6V ~ 4.2V |  |
| Working current |  | ≤ 80 mA |  |
| Power supply |  | A rechargeable lithium battery (3.7 V) (The red wire on the battery denotes anode, the black wire on the battery denotes cathode). |  |
| Battery life |  | Charge and discharge: no less than 500 times. |  |
| Adapter specification |  | Output voltage: DC 5V Output current: 1000 mA |  |
|  | Dimension and Weight |  |  |
| Dimension |  | 46(L) x 40(W) x 29(H) mm |  |
| Weight |  | About 35 g (including a lithium battery) |  |
|  | Display |  |  |
| Display Size |  | 1.25inch |  |

| Appendix 1 |  |  |
| --- | --- | --- |
| State | Sound prompt condition delay | Sound prompt signal generation delay |
| Low voltage sound prompt | 1s | 20ms |
| SpO2 sound prompt | 330ms | 20ms |
| Pulse rate sound prompt | 330ms | 20ms |
| Probe error sound prompt | 16ms | 20ms |